## Slide 1: Outline: RISING – active stopper
Lecturer: Hans-Jürgen Wollersheim
e-mail: h.j.wollersheim@gsi.de
web-page: https://web-docs.gsi.de/~wolle/ and click on
measurements with double-sided Si-strip detector
mesytec and multichannel systems electronics
241Am and 207Bi source
implantation detector for RISING
experimental results

## Slide 2: Implantation detector as active stopper 2006
Implantation-decay correlations with large background
(half lifes similar to the implantation rate):
implantation-decay time correlation: active catcher
implantation-decay position correlation: granularity
implantation of several ions: thickness and area
energy of the implanted ion and the emitted b
Active catcher for implantation-decay correlations
3 double-sided silicon-strip detectors
- surface 5x5 cm2
- thickness 1 mm
- 2 x 16 3.125 mm strips
- manufactured by MICRON
Rakesh Kumar, P. Doornenbal, I. Kojouharov, W. Prokopowicz, H. Schaffner, H.J. Wollersheim NIM A 598 (2009) 754

## Slide 3: Measurements with a double-sided Si-strip detector 2006
32
Front Junction Side
Rear Ohmic Side
CAEN
Micron Semiconductor
Nº Junction Elements: 16
Nº Junction Elements: 16
Element Length: 49.5 mm
Element Pitch: 3.1 mm
Element width: 3.0 mm
Active Area: 50x50 mm2
Thickness: 1000 µm
Price: 5600 €
MPR-32
Charge Sensitive Preamplifier
32 channel compact module
Sensitivity switch, factor 5
Bias voltage up to ±400V
Price: 2790 €
STM-16
16 fold shaper
16 channel NIM module
shaper amplifier
timing filter amplifier
leading edge discriminator
Price: 2x 3415 €
MRC-1
rc master controller
for STM-16
Price: 2200 €
ADC V785AF
32 channel
Price: 5094 €
Total cost 22,514.- €

## Slide 4: Energy resolution with 241Am source measurement in vacuum
Front Junction Side
Rear Ohmic Side
Low energy peak from gap events at about ½ the full pulse height
C.Wrede et al. NIM B204 (2003), 619
MICRON #2215-17
Voltage: 200V
241Am Eα=5.486 MeV
range ~28 µm

## Slide 5: Strip multiplicity with 241Am source
Front Junction Side
Rear Ohmic Side
MICRON #2215-17
Voltage: 40V
below full depletion
measurement in vacuum
MICRON #2215-17
Voltage: 200V
full depletion voltage
measurement in vacuum
241Am Eα=5.486 MeV
range ~28 µm
Front Junction Side
Rear Ohmic Side

## Slide 6: Two dimensional position spectra
MICRON #2243-5 Voltage: 40V, measurement in vacuum
241Am source centered
241Am source left
241Am source left, strip-multiplicity=1
241Am source centered, strip-multiplicity=1

## Slide 7: Energy resolution with 207Bi source measurement in vacuum
Front Junction Side
Rear Ohmic Side
MICRON #2512-17
Voltage: 200V
207Bi E=482, 976 keV
range 0.94, 2.31 mm (e-e- interaction)
experimental set-up

## Slide 8: Energy resolution with 207Bi source measurement in vacuum
Front Junction Side
Rear Ohmic Side
MICRON #2512-17
Voltage: 200V
207Bi E=482, 976 keV
range 0.94, 2.31 mm (e-e- interaction)
experimental set-up

## Slide 9: Energy resolution with 207Bi source measurement in vacuum
Front Junction Side
Rear Ohmic Side
MICRON #2512-17 (full lines)
#2243-5 (dashed lines)
Voltage: 200V
207Bi E=482, 976 keV
range 0.94, 2.31 mm (e-e- interaction)
experimental set-up

## Slide 10: Energy resolution with 207Bi source measurement in vacuum and dry N2
Front Junction Side
Rear Ohmic Side
MICRON #2243-5
Voltage: 200V
conclusion: measurement in dry N2
207Bi E=482, 976 keV
range 0.94, 2.31 mm (e-e- interaction)
experimental set-up

## Slide 11: Energy resolution of the DSSSD
207Bi Ee=976 keV ; 241Am Eα=5.486 MeV
experimental set-up
MICRON | ∆E (241Am) vacuum | ∆E(207Bi) vacuum | ∆E(207Bi) dry nitrogen
#2243-5 | N: 31.3 keV | N: 16.2 keV P: 33.3 keV | N: 16.0 keV P: 32.5 keV
#2243-4 | N: 30.2 keV | N: 18.5 keV | 
#2243-3 | N: 34.0 keV | N: 18.2 keV | 
#2243-2 | N: 35.7 keV | N: 14.5 keV P: 27.0 keV | 
#2512-17 | N: 27.4 keV P: 29.7 keV | N: 14.8 keV P: 18.8 keV | 
Front Junction Side
Rear Ohmic Side
∆E=15.4 keV
∆E=32.5 keV

## Slide 12: RISING: Test of the active stopper
γ-energy [keV] | e--energy
569.6 | 481.7 [K]
 | 553.8-556.7 [L]
 | 565.8-567.2 [M]
1063.7 | 975.7 [K]
 | 1047.8-1050.6 [L]
 | 1059.8-1061.2 [M]
207Bi emits gamma rays and electrons

## Slide 13: Energy threshold of the DSSSD
207Bi Ee=482, 976 keV
experimental set-up

## Slide 14: Measurements with a double-sided Si-strip detector 2006
32
Front Junction Side
Rear Ohmic Side
CAEN
Micron Semiconductor
Nº Junction Elements: 16
Nº Junction Elements: 16
Element Length: 49.5 mm
Element Pitch: 3.1 mm
Element width: 3.0 mm
Active Area: 50x50 mm2
Thickness: 1000 µm
Price: 5600 €
CPA-16
Charge Sensitive Preamplifier
16 channel compact module
2 output stages with different gains
Bias voltage up to ±500V
Price: 2x 2250 €
Amplifier N 568BC
16 fold shaper
Price: 2x 3481 €
ADC V785AF
32 channel
Price: 2x 5094 €
CAEN
Total cost 27,250.- € (discriminator not included)

## Slide 15: Energy resolution with 207Bi source measurement with Mesytec and Multichannel Systems
Front Junction Side
MICRON #2243-5
Voltage: 200V
measurement in vacuum
ORTEC 572
shaping time 0.5 μs ∆E=122 keV
1.0 μs ∆E=112 keV
2.0 μs ∆E=103 keV
experimental set-up

## Slide 16: Energy resolution with 207Bi source measurement with Multichannel Systems
MICRON #2243-5
Voltage: 200V
measurement in vacuum
experimental set-up
ORTEC 573
CAEN N568BC

## Slide 17: Implantation detector as active stopper 2006
Implantation-decay correlations with large background
(half lifes similar to the implantation rate):
implantation-decay time correlation: active catcher
implantation-decay position correlation: granularity
implantation of several ions: thickness and area
energy of the implanted ion and the emitted b
Active catcher for implantation-decay correlations
3 double-sided silicon-strip detectors
- surface 5x5 cm2
- thickness 1 mm
- 2 x 16 3.125 mm strips
- manufactured by MICRON
Rakesh Kumar, P. Doornenbal, I. Kojouharov, W. Prokopowicz, H. Schaffner, H.J. Wollersheim NIM A 598 (2009) 754

## Slide 18: Double-sided silicon-strip detector DSSSD
thickness sufficient for HI-implantation
but range of β-particles larger than 1mm
therefore part of the kinetic energy is measured
- surface 5x5 cm2
- thickness 1 mm
- 2 x 16 3.125 mm strips
- manufactured by MICRON
Monte-Carlo Simulation with GEANT4

## Slide 19: Monte-Carlo simulation with GEANT4
Simulated energy spectrum of β-particles emitted from fragments implanted uniformly (solid line) and exactly in the centre (dashed line) in the middle of a DSSSD.
Calculated β-detection efficiency as a function of the DSSSD threshold for the two considered implantation scenarios
Detection threshold should be less than 100 keV

## Slide 20: Rare ISotope INvestigation at GSI spectroscopy of stopped beams
identified projectile fragments
DSSSD for HI and β
n - det.
g - det.
g - det.
projectile implants
projectile decays
implantation ↔ range focusing
position correlation ↔ high granularity
time correlation ↔ dedicated electronics

## Slide 21: Implantation detector
double sided silicon strip detector
active area 50x50 mm2
thickness 1.0 mm
16 strips in x- and y-direction

## Slide 22: Chamber for active stopper measurement with dry N2
Pertinax
phenolic-formaldehyd cellulose-paper
PF CP 2061
result:
6mm Pertinax ≈ 2mm Al
2mm Pertinax for active stopper chamber

## Slide 23: Stopped RISING array @ GSI: 15x7 element Cluster with passive Cu stopper

## Slide 24: Stopped RISING array @ GSI: 15x7 element Cluster with DSSD

## Slide 25: Chamber for active stopper
Thin black Pocalon C foil 20um

## Slide 26: Stopped RISING array @ GSI: 15x7 element Cluster with DSSD

## Slide 27: RISING set-up with stopped beams
RISING Ge
Cluster Array
MUSIC
MUSIC
Sci41
MW41
MW42
degrader
beam

## Slide 28: RISING with stopped beams
no. of Clusters | angle | distance to target
5 | 500 | 210+10mm
5 | 900 | 210+20mm
5 | 1300 | 210+10mm
beam
photopeak efficiency:
18.2%
J. Simpson CCLRC Daresbury

## Slide 29: RISING set-up with stopped beams

## Slide 30: Count rate limitations with active stopper
3 x 16 x 16 = 3 x 256 = 768 total pixels.
  Assume upper limit for β-half-life of ~30 seconds
  Each pixel hit every 5 half-lives (150 s)
Max. rate of ~768/150 = 5 per sec ( = 50 per 10s spill).
Rate increases directly with decreasing half-life
(e.g., T1/2 = 10 seconds -> 150 per 10 s spill cycle)
Dual gain pre-amps on DSSSD to get energies of
implanted ion and β-particle
All events time stamped with MHz clock.

## Slide 31: Fragment separator FRS
Wedge-shaped
Degrader
19Ne at 600AMeV:
Phase-space imaging of differently shaped degraders within the achromatic ion-optical system. The results for a homogeneous, an achromatic, and a monoenergetic degrader are given. All degraders have the same thickness on the optical axis (d/r=0.5)
Primary
Beam
Beam & All
Fragmentation
Products
Momentum
Selection
Spacial
Dispersion
Secondary
Beam
Isotope
Selection

## Slide 32: Fragment separator FRS 40Ar 50 MeV/u + Ta (100 μm), wedge shaped Al (200 μm) degrader
0.39 mrad
1.66mrad

## Slide 33: Chromatic aberration
When different colors of light propagate at different speeds in a medium, the refractive index is wavelength dependent. This phenomenon is known as dispersion.
Longitudinal (axial) chromatic aberration: Transverse (lateral) chromatic aberration:
The focal planes of the various colors do The size of the image varies from one
not coincide. color to the next.

## Slide 34: Implantation range
Estimated implanted isotopes for a setting centered on 192W
in 1 mm thickness silicon with a monoenergetic degrader at S2
Silicium thickness (mm)
charge states

## Slide 35: Estimated implanted isotopes for a setting centered on 192W in 1mm thickness silicon with a monoenergetic degrader at S2
Silicium thickness (mm)
charge states
Setting centered on 198Ir
produced implanted

## Slide 36: Experimental results conversion electrons
Zs. Podolyak, Phys. Lett B672 (2009), 116
proton-hole excitation in the closed shell nucleus 205Au

## Slide 37: Experimental results beta decay
F. Molina, Phys. Rev. C91 (2015) 014301
Tz-1 → Tz=0 GT transition (β-decay)